## Forms
Does the site have an individual identified as the POC for instrument reprocessing? | Does the POC have record of training/competency in instrument reprocessing (reviewer should visually verify)? | Does the site have an instrument reprocessing SOP available (reviewer should visually verify)? | Comments and Recommendations (Leadership) | Were the instruments rinsed or wiped off with a damp gauze to remove gross contamination immediately after the case? | Is there a process to flag and communicate instrument problems (e.g., dull scissors or non-locking clamp)? | Was prolonged instrument soaking (>10 minutes) avoided? (soaking instruments prior to cleaning is not recommended, however, if it is done, it should be <10 minutes | If soaking was used, was the use of saline, chlorine, or disinfectant avoided? (plain water or detergent (soap) are ok; please note that soaking in any chlorine solution, a common procedure in the past, should not be done) | Were instruments left open (i.e., jaws not clamped shut) during and after pre-cleaning? (If instruments are clamped shut, closed surfaces and hinges cannot be cleaned which allows organic material to dry and become stuck) | Were instruments transported from point of use to cleaning area in a leak-proof, puncture-proof, and covered container? | Were instruments kept moist the entire time between end of case and cleaning? | Comments and Recommendations (Pre-cleaning) | Was appropriate PPE (hair cover, eye/mouth protection, impermeable gown, gloves, closed shoes) worn during cleaning? | Was a clean container filled with an appropriate amount of water and detergent to allow complete immersion of the instruments? | Were the instruments cleaned using a soft (nylon) bristle brush under the surface of the water (to avoid producing aerosols)? | Were all hinged instruments opened during cleaning? | Were instruments immersed or rinsed thoroughly with clean water after cleaning with the brush? (soft (e.g., distilled) water preferable if available to prevent deposits and buildup on instruments) | Were instruments closely inspected after cleaning, prior to drying, to make sure all visible contamination had been removed? | Were instruments cleaned in an area that is adequately illuminated? (the inspection area should have adequate illumination (natural light may be adequate,) to see small amounts of contamination) | Were instruments air-dried or dried by hand with a clean, non-linting cloth? | Was the overall flow of instruments in the cleaning area unidirectional from “dirty” to “clean” | If both reusable and single use instruments are used, is there an established way these are separated and distinguished from one another? (Assessor should ask for a description of this process. Single use instruments are typically marked by the words “single use” or the number '2' with a strike through it). | Comments and Recommendations (Cleaning) | Were instruments checked for proper functioning, and lubricated as necessary to allow free movement at hinges, prior to packaging? | Are kits assembled to ensure completeness and standard contents? (sites often have an instrument list posted in the packaging area as a job aid) | Is the packaging material permeable to the sterilizing agent (e.g., steam)? Packaging material examples include: double layer of lint-free dense woven cloth (≥140 threads/inch2 ), unwoven cloth permeable to steam (i.e., sterilization wrap), surgical-grade paper (not recycled paper or newspaper), or rigid reusable sterilization containers. | Does packaging securely enclose items to be sterilized? | Does packaging maintain sterility of contents until opened (doesn’t easily puncture or allow contamination inside)? | Does every package have an external chemical indicator? (external chemical indicators show when a package has been exposed to one or more sterilization parameters such as temperature, pressure, time; external indicators are used to differentiate processed and non-processed items) | Does every package contain an internal chemical indicator? (internal chemical indicators show that the instruments themselves, not just the outside of the package, were exposed to one or more sterilization parameters) indicators are used to differentiate processed and non-processed items) | Is every package labeled with sterilization date and/or expiry date, contents (e.g., “vmmc kit”), and load number? | Comments and Recommendations (Packaging) | Is steam sterilization in use? If yes, answer the following: | Has the person responsible for operating the autoclave received training on how to do so? (Sometimes facility staff, not VMMC staff, run the autoclave. This question explores presence and type of training they received. Please check documentation (e.g., training certificate) if available.) | Is cycle temperature either pre-set by load type or, if manually set, to at least 121⁰C ? Cycle times will vary (3-18 minutes) depending on temperature (121⁰C - 135⁰C), please refer to manufacturer’s instructions for use, if available, to check that manually set time is correct for temp. | Are the following cycle parameters recorded in a logbook? | Is the autoclave loaded in a manner to ensure steam contact and penetration? (not overloaded, packages away from chamber walls, rigid containers not stacked unless validated by manufacturer for this configuration, steri-peel packages on edge with multiple packages placed paper-to-plastic, non perforated containers placed on edge) | Correct cycle parameters met? (if pre-set parameters, could just mean the ‘failure’ indicator did not light up) indicators are used to differentiate processed and non-processed items) | External indicators appropriately changed? | No visible signs of moisture? (no dampness, droplets, or puddles on or within a package) | No visible signs of compromised package integrity? | Bowie-Dick test (measures air removal in pre-vacuum autoclaves, not used in gravity displacement autoclaves; ideally performed daily) | Biological indicators (please make note how often these are used (ideally should be daily, but at least weekly) | Is distilled water being used in the autoclave? (most autoclave manufacturers recommend use of distilled water to avoid mineral deposits which can decrease useable life of the autoclave) | Over the past 6 months has the distilled water inventory been maintained to ensure uninterrupted services? | Is there documentation of at least annual routine maintenance performed by a technician? (if “no” then please ask follow-up question about why this was not done) | Was the overall flow of instruments in the sterilization area unidirectional from “non-sterile” to “sterile”? | Are all biologic and chemical indicators within their expiration dates? | Over the past 6 months has the reprocessing equipment and supplied inventory been maintained to ensure uninterrupted services? | Over the past 6 months has the reprocessing equipment and supplied inventory been maintained to ensure uninterrupted services?1 | Comments and Recommendations (Sterilisation) | Is the portion of the transport vehicle where instrument containers are placed completely enclosed? | Are the containers for sterile and non-sterile instruments clearly marked? | Is a system in place to track movement of instrument kits between sites, and know where they are at all times? (for example, to find where all kits sterilized on x date in cycle number y are located; if one kit from a specific load is noted to have problems, it will be important to immediately identify all the other kits that were in the same load). | Comments and Recommendations (Transportation) | Are freshly sterilized instruments allowed to cool to room temperature before they are used for surgery? | Is a practice of first in first out used to ensure that older kits (by date of sterilization) are used first? | Is there an inventory monitoring system that tracks the movement of kits in and out of the storage or sterilization area? | Comments and Recommendations (Storage)
yes | yes | yes | Leadership is strong | yes | yes | yes | yes | yes | yes | yes | Pre Cleaning was done at a high standard | no | yes | yes | yes | yes | no | yes | yes | yes | no | Needs to Improve Didnt cover hair | yes | yes | yes | no | yes | yes | yes | yes | This was done to the highest standard | yes | yes | no | Date_of_sterilization Cycle_number Cycle_length Cycle_temperature external_indicator_change | yes | yes | yes | yes | yes | yes | no | yes | yes | yes | yes | yes | yes | yes | Didnt reach the required temperature | yes | yes | yes | This was done very well | yes | yes | yes | Strong storage mechanism